# Tribological Study of Bioinert Zirconia-Toughened Alumina Used in Total Hip Joint Implants

Source: Research Square preprint, posted 2022-04-26. DOI: 10.21203/rs.3.rs-1584398.

URL: https://www.researchsquare.com/article/rs-1584398/v1

Authors: SHARANRAJ V, C. M. Ramesha, T. Anil Kumar, Abhilash S, and others.

Licensed under a Creative Commons Attribution (CC BY 4.0) license (https://creativecommons.org/licenses/by/4.0/). This preprint has not been peer reviewed by a journal. This excerpt (abstract and introduction) is reproduced verbatim as raw material for a before/after editing sample.

## Abstract

Zirconia-toughened alumina (ZTA) is an advanced ceramic bioinert biomaterial comprising 80% alumina and 20% yttria-stabilized zirconia. It is used in several medical applications, including  total hip joint implants, owing to its unique characteristics such as its low interaction with the surrounding tissue and good biocompatibility. This paper presents the tribological (wear) behavior of ZTA pins and plates when subjected to wear tests. A Ducom Linear Reciprocating Tribometer was used for evaluating the wear characteristics of the ZTA pin and plate specimens. The average wear rate of the specimens was low at 6.2498  10-9 g/(Nm). Their coefficient of friction was 0.333 for a frictional force of 6.666 N and operating temperature of 37 ± 0.2 C. These values are within the admissible limits as per the ASTM standards for implant materials. Hence, ZTA is a suitable biomaterial for total hip joint implants and other similar medical applications.

Keywords: Zirconia, Alumina, Wear resistance, Composites, Biomedical applications

## 1. Introduction

Zirconia-toughened alumina (ZTA) is an advanced ceramic bioinert biomaterial comprising alumina and zirconia. It consists of 80% alumina, which is the primary phase, and 20% yttria-stabilized zirconia, which is the secondary phase. This composite ceramic material combines the advantageous properties of both monolithic alumina and zirconia. As a result, it exhibits high strength and toughness compared with conventional alumina and zirconia. In addition, ZTA composites show better mechanical properties compared with monolithic alumina or stabilized zirconia alone [1]. This can be attributed to several mechanisms such as blocked crack growth, grain shape control, control of the phase transformation of the zirconia particles, and the strengthening of the alumina phase by optimizing its grain size and introducing various additions.

While alumina exhibits high hardness and good wear resistance, its high brittleness results in high fracture risk, leading to design restrictions. The hardness of zirconia is half of that of alumina; however, its fracture resistance is greater owing to transformation toughening. Hence, its overall toughness and bending strength are comparatively higher than those of alumina. In addition, zirconia exhibits a metastable form phase transition, which affects its overall stability [2]. The poor thermal conductivity of zirconia also contributes to this phenomenon. Therefore, the ideal ceramic would combine the best properties of both zirconia and alumina: one such material is zirconia-toughened alumina.

A biomaterial is a natural or synthetic material used for medical applications because it is compatible with living tissues. The important factors considered for a biomaterial are its physical properties, corrosion characteristics in the tissue environment, and surface configuration. Based on the biological interactions of biomaterials, they can be categorized into three types: bioinert materials, bioactive materials, and bioresorbable materials. Biomaterials may be metals and metallic alloys, ceramics, polymers, or natural materials. The function of biomaterials ranges from being inert in the body to being "bioactive" and helping with regeneration [3].

Biomaterials are widely used for the restoration of damaged body parts through interactions with the living system. Biomaterial devices used in orthopedics are known as implants. Biomaterial-based implants used in the human body include artificial heart valves; blood vessel stents; shoulder, knee, and hip implants; elbow replacement implants; and dental implants.

A biomaterial should have the following desired properties [4]:

Mechanical properties: The material should have high strength so that it can withstand high loads. In addition, its modulus of elasticity should be low. This would increase the service life of the implants and prevent the necessity of revision surgery owing to the loosening of the implant.

Biocompatibility: The material should not cause any harm to the body and should be compatible with living systems.

Wear resistance: The material must possess high wear resistance and a low coefficient of friction when sliding against body tissues. A low wear resistance or high coefficient of friction can lead to loosening of the implant.

Corrosion resistance: A biomaterial must possess high corrosion resistance. Materials with low corrosion resistance can release metal ions within the body, resulting in toxic reactions.

Osteoconductivity: The biomaterial must allow bone growth on its surface.

Osseointegration: The properties of the material should be such that both functional and structural connections form between the surface of the load-carrying implant and the living bone. The roughness, chemistry, and surface characteristics of the implant material play a major role in osseointegration.

Nontoxicity: The implant material must be nontoxic, that is, it must not result in toxic reactions within the body and should not alter the DNA (genotoxic) or cause damage to the individual cells (cytotoxic).

Fatigue life: The material must have high resistance to fatigue so that implant failure by fatigue fracture can be prevented.

Bioinertness: The material must not interact when introduced within biological tissue. This helps in the formation of bone on the implant surface and aids the ion exchange with the surrounding tissues.

Biotolerance: The material must not be rejected when implanted into living tissue.

Biomimetics: The material must be designed to mimic specific biological processes, so that it helps optimize the regenerative response within the biological environment.

Total hip joint implant replacement is an unavoidable orthopedic procedure for arthritis patients seeking to improve their quality of life [5]. Artificial joints that replace the damaged joint are gaining popularity. The demand for hip joint replacement in orthopedic surgery is high, but making an implant last remains a challenge.

Standard total hip joint implants are made up of four components [6]. The stem is the part inserted into the femur, or thighbone. The cup is the part inserted into the pelvic bone. The ball is the part fitted onto the end of the stem. And the liner is the part inserted into the cup; it acts as the new cartilage.

Tribology is "the science and technology of interacting surfaces in relative motion and of related subjects and practices." It generally includes technologically relevant phenomena such as friction, wear, and lubrication [7, 8]. Wear is the progressive loss of material from surfaces in contact, which occurs because of friction. It is a slow, continuous process of particle removal from one or more materials. Wear is not an intrinsic property of a material but rather a complex response of a material induced by the system characteristics, such as the load, speed, temperature, hardness, surface finish, and presence of foreign materials [9]. The processes by which material experiences wear are known as its wear mechanisms. These mechanisms describe how the articulating surfaces behave under different tribological conditions [10]. The types of wear experienced in a material may be abrasion, adhesion, third-body, fatigue, and corrosion [11].

The parameters affecting wear include the coefficient of friction, frictional force, and operating temperature.